## Slide 1: Final Demo Meeting WP7 summary and results 25-Jan-2023
P.Finocchiaro

## Slide 2
WP7 Waste Monitoring Grid
  T7.1 Assess the feasibility of the SiLiF (neutron) and SciFi (gamma) detectors to monitor the different types of nuclear waste
  T7.2 Build, test and characterize the detectors’ behavior with respect to different types of nuclear waste and containers
  T7.3 Characterisation of the system with laboratory tests in order to prepare the final demonstration
  T7.4 Performance evaluation of alternative technologies (Timepix network)
Tasks
100%
100%
100%
100%

## Slide 3
D7.1 | INFN - Design of the integrated prototype | Report of the integrated system that will be used during the final demonstration | M10 | 
D7.2 | INFN - Integrated monitoring grid system | Deployable monitoring grid prototype | M24 | 
D7.3 | INFN - Report on the monitoring grid system laboratory characterisation | Description of the characteristics of the system with description of the laboratory characterisation | M28 M36 | 
D7.4 | CAEN - The prototype of the monitoring grid system | System ready for the deployment in a real environment | M28 M36 | 
D7.5 | CTU - Report on the characterisation of alternative solutions for long-term monitoring system | Description of the characteristics of the system based on Timepix detectors and main laboratory results | M27 M36 | 
D10.1 | INFN - The waste monitoring grid system in a nuclear waste repository | Report of the characterization of the waste monitoring grid in real environments | M34 M41 | 
Mar 2020
May 2021
May 2022
May 2022
May 2022
Oct 2022
Deliverables

## Slide 4
test @ENEA
Waste assessment completed
Brief report with the list of candidate waste(s) and the related detector performance expected.
12
May
2020
13
Detector assembled
Construction of Fiber and SiLiF detector samples ready for laboratory tests
May
2021
14
Waste monitoring grid system
Detectors ready for in field tests
22
Detector deployment
Deployment of the integrated monitoring system in a real environment
choose config
build & lab test
deploy
final demo
T7.1
T7.2
T7.3
(WP10)
WP10
May
2022
Dec
2022
test
deploy
test
Jul
2022
Milestones

## Slide 5
SciFi gamma ray detectors
SiLiF thermal neutron detectors
PET moderators
what was built

## Slide 6
SiLiF neutron detector: Silicon + 6LiF
10 x 10 x 10 cm3
PET moderator to
slow down neutrons
flat efficiency after moderation
half moderator
4x4 moderators
6LiF on C-fiber
Si diode
6LiF on C-fiber
Si
6LiF
C
C
6LiF
6LiF evaporation
on C-fiber
GEANT4
simulation

## Slide 7
SiLiF neutron detectors built and tested
AmBe neutron source
2.2·106 n/s
test in front position
FLUKA simulation
background
neutrons

## Slide 8
SiLiF neutron detectors: 5% efficiency measured
MICADO detectors: 2x17 µm 6LiF
reference detector: 1.6 µm 6LiF
calibrated at PTB metrology
thermal eff. = 0.41% ± 0.001%
AmBe neutron source
2.2·106 n/s
test in inside position
FLUKA simulation
thermal eff. = 5.0%

## Slide 9
SiLiF neutron detectors: gamma/neutron rejection
measurement of background
with upside-down converters
knowledge of pure neutron
and pure background to calculate gamma/neutron discrimination power
gammas with E ≤ 4.4 MeV
gammas with E ≤ 2.2 MeV
target: 10-5 with Eg ≤ 1.3 MeV (“optimal” according to IAEA)
obtained: 10-10 with Eg ≤ 2.2 MeV (SiLiF with threshold at 1.5 MeV)

## Slide 10
16 SiLiF neutron detectors: sensitivity tested with fast neutrons
test in top position
expected 2.09 n/s/cm2
measured 1.92 n/s/cm2
simulated overall detection efficiency
energy dependent detection efficiency
simulated neutron flux in det05

## Slide 11
SciFi gamma detector built and tested with pointlike sources
SiPM 3x3 mm2
(ON Semiconductor)
holder
holder
Aluminum pipe
mechanical assembling
operating principle
simulated detection efficiency
(and interpolation)
average measured efficiency 1.65% ± 0.18%
simulated efficiency 1.82%
with 137Cs source
SiPM sensitive to
single photons

## Slide 12
SciFi gamma detectors tested with a drum-like AmBe source box
emitted gamma spectrum
measured with LaBr3(Ce)
results agree with simulations

## Slide 13
10x Timepix3 + Katherine R/O
  Efficient particle discrimination, spectrometric information in parallel with ns resolution timing information
  Si sensors of 100, 300 & 500 µm,
  Equipped with neutron converters
  CdTe sensors of 1 mm
  Operated simultaneously by specific SW
Thermal neutrons Neutrons from AmBe Neutrons from Cf Neutron spectrum in CdTe 137Cs spectrum
Timepix3 detector network tested
662 keV

## Slide 14
Timepix3 software developed, multiple device control

## Slide 15
Tests at ENEA C43 lab
  Timepix3 detectors (all 4 sensor types)
α particles
from Rn decay
137Cs
662 keV

## Slide 16
Monitoring configuration: flexible, scalable
Possible arrangements
on the drum
on the platform
SciFi (gammas)
SiLiF (neutrons)
SiLiF
SciFi
Timepix3 (n and g)
TimePix

## Slide 17
neutrons
gammas
DAQ CAEN ( DigiWaste)
Detectors and support structures
Integration and tests of detectors with CAEN electronics at INFN-LNS
Tested with special sources at ENEA
flexible support structure for quick and easy reassembling using Bosch Rexroth aluminum profiles
sample data
SiPM sensitive to
single photons

## Slide 18
The simple & lightweight support structure with four SiLiF+SciFi detectors
.
Test at ENEA C43-lab 27-30 May 2022
The detectors in position to accommodate a large 220 liter mock-up drum
Test with gamma and neutron sources in the mock-up drum
137Cs
2.33 MBq
240Pu
4×103 n/s
the data acquisition system
1
2
3
4

## Slide 19
.
SciFi gamma detectors
1
2
3
4
SiLiF neutron detectors
2
4

## Slide 20
 | gamma/s | stat error | background | stat error
SciFi 1 | 5.28 | 0.06 | 2.76 | 0.02
SciFi 2 | 9.57 | 0.07 | 1.85 | 0.02
SciFi 3 | 26.14 | 0.12 | 2.15 | 0.02
SciFi 4 | 6.18 | 0.06 | 2.64 | 0.02
 | n/s | stat error
SiLiF4 | 0.0321 | 0.0007
SiLiF6 | 0.0417 | 0.0008
SiLiF2 | 0.0363 | 0.0007
SiLiF15 | 0.0372 | 0.0007
18 hour test
1/2 hour test
240Pu neutron source (4·103 n/s)
2.33 MBq 137Cs gamma source

## Slide 21
from ENEA Casaccia to ORANO La Hague

## Slide 22
Test at ORANO La Hague on four drums (with Pu) 25-28 Jul 2022

## Slide 23
Test at ORANO La Hague on four drums (with Pu)
SciFi gamma detectors
gamma source

## Slide 24
Test at ORANO La Hague on four drums (with Pu)
SiLiF neutron detectors
RP people measured and declared << 1 n/s
(basically zero)

## Slide 25
from ORANO La Hague back to ENEA Casaccia

## Slide 26
Test at ENEA C43-lab 12-13 Dec 2022
mock-up C
too noisy electrical environment
gamma
neutron
1000× for SiLiF?!?!
ineffective threshold
but results in agreement with source position and activity

## Slide 27
from ENEA C43-lab to Nucleco
ENEA
C43

## Slide 28
Tests at Nucleco:
14-15 Dec 2022
15-22 Dec 2022
23 Dec 2022 to 26 Jan 2023

## Slide 29
neutrons at Nucleco 21-Dec-22 to 11-Jan-23
stable SiLiF data rates
N-01 has no counts: wrong threshold
E-12 and W-06 have few counts: wrong threshold?

## Slide 30
gammas at Nucleco 21-Dec-22 to 11-Jan-23
stable SciFi data rates
slight variations due to temperature changes
some slight change due to local operations?

## Slide 31
gammas at Nucleco 21-Dec-22 to 11-Jan-23
fluctuations within 1 sigma
people from CAEN operating on the system

## Slide 32
Sensors-21-02630
Instruments-05-00019
SiLiF & SciFi for MICADO [talk]
Timepix3 for MICADO [e-poster]
I rivelatori del progetto MICADO per il monitoraggio dei rifiuti radioattivi
Martina Giuffrida, Master Thesis in Physics
7-Jun-21 110/110 cum laude
Rivelatori di gamma e neutroni per il monitoraggio di rifiuti radioattivi nel progetto MICADO [talk]
Activity Report 2020
Publications
peer reviewed journals
conferences & workshops
other papers
theses
Troyes (France), 04-08 July 2022
[2 posters]
Symposium on International Safeguards
31 Oct to 4 Nov 2022 [talk]
General Conference, 26-30 Sep 2022 [talk]
P. Finocchiaro, Providing better radiation monitoring for nuclear waste, Res. Outr. 131(2022)70
Technology Park Workshop
Rome 20-21 Jan 2023
Radioactive Waste monitoring [invited talk]

## Slide 33
Conclusions
  The MICADO project has been a wonderful experience
  We got in touch and collaborated with many people
  Within WP7 we had the chance of studying and improving our technologies for gamma and neutron detection
  These detectors proved to be feasible for a distributed radwaste monitoring application
  I thank all of the MICADO people, but in particular Massimo Morichi, Erica Fanchini, Luigi Cosentino, Alfio Pappalardo, Fabio Longhitano, Gaetano Elio Poma, Simone Amaducci, Sergio Lo Meo, Quentin Ducasse, Bartolomej Biskup, Lionel Tondut, Christopher Helbert, Bernard Dumercq, Luigi Lepore, Giada Gandolfo, Giuseppe Marzo, Nicolò Ratto, Nadia Cherubini, and last but not least my dear friend Marco Ripani
  ...and I apologize if I am (unfortunately) forgetting someone